## Traceability Matrix
Risk ID | Title | Component ID | Component | Subsystem | Failure mode ID | Failure mode | Cause | Hazard | Hazardous situation ID | Hazardous situation | Harm ID | Harm | Detection before control | S pre | P pre | D pre | RPN pre | Region pre | Priority pre | Probability basis | Similar incidents | Risk control measures | Requirements | Verification evidence | S post | P post | D post | RPN post | Region post | Priority post | RPN reduction pct | Disposition | Design specification references | Standard clause references | Incident evidence
RISK_ALM_01 | Speaker failure | CMP_ALM_01 | Audible and Vibratory Alarm Hardware | Alarm | FM_ALM_01 | Speaker fails and the audible alarm is silent | Open circuit of the speaker element or a blocked sound port | Loss of alarm information | HS_ALM_01 | An occlusion or low battery alarm is shown on the display but is not heard during sleep | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 150 | RCM_ALM_01 Speaker self test by current sensing with redundant vibration annunciation [protective measure, verification PASS, credited P2 D2] | REQ_ALM_001 | VER_ALM_001 PASS TR_0120 | 4 | 1 | 3 | 12 | ACCEPTABLE | MEDIUM | 80 | Residual risk acceptable | HDS_002 section 7.1 (DFC_ALM_01); REQ_ALM_001 at HDS_002 section 7.1 | IEC 60601 1 8 cl 6.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_003614; MDR_2022_008328; MDR_2023_008370; MDR_2024_003265; MDR_2025_006760
RISK_ALM_02 | Alarm not perceived | CMP_ALM_01 | Audible and Vibratory Alarm Hardware | Alarm | FM_ALM_02 | Alarm volume is insufficient for the user to perceive the alarm | Low volume setting or the pump is muffled under clothing and bedding | Loss of alarm information | HS_ALM_02 | A no delivery alarm sounds but is not heard and delivery remains stopped | HARM_HYPER_SERIOUS | Sustained hyperglycemia requiring medical intervention | User observation | 3 | 4 | 4 | 48 | REVIEW | HIGH | engineering estimate | 563 | RCM_ALM_02 Escalating alarm volume with repeated vibration until acknowledged [protective measure, verification PASS, credited P2] | REQ_ALM_002 | VER_ALM_002 PASS TR_0121 | 3 | 2 | 4 | 24 | ACCEPTABLE | MEDIUM | 50 | Residual risk acceptable | HDS_002 section 7.1 (DFC_ALM_02); REQ_ALM_002 at HDS_002 section 7.1 | IEC 60601 1 8 cl 6.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_008069; MDR_2022_006288; MDR_2023_007790; MDR_2023_009109; MDR_2023_009651
RISK_CGM_01 | False high glucose drives automated correction | CMP_CGM_01 | Glucose Sensor Data Interface | Connectivity | FM_CGM_01 | Falsely high glucose reading is used for automated correction | Sensor compression artifact or calibration error passes without a plausibility check | Overdelivery of insulin | HS_CGM_01 | Automated correction delivers insulin against a false high reading | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 3 | 5 | 75 | UNACCEPTABLE | HIGH | engineering estimate | 100 | RCM_CGM_01 Rate of change plausibility filter with a cap on automated correction doses [protective measure, verification PASS, credited P2]; RCM_IFU_06 Labeling that instructs fingerstick confirmation before large corrections [information for safety, verification PASS, credited P1] | REQ_CGM_001; REQ_LBL_006 | VER_CGM_001 PASS TR_0150; VER_LBL_006 PASS TR_0155 | 5 | 1 | 5 | 25 | REVIEW | MEDIUM | 66.7 | Benefit risk analysis required | UCS_004 section 4.2 (DFC_CGM_01); REQ_CGM_001 at UCS_004 section 4.2; REQ_LBL_006 at UCS_004 section 5.1 | IEC 62304 cl 7.2; IEC 62366 1 cl 5.9; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2022_009287; MDR_2023_005338; MDR_2024_002356; MDR_2025_005294; MDR_2025_007146
RISK_CHG_01 | Charger isolation failure | CMP_CHG_01 | Charging Circuit | Power | FM_CHG_01 | Isolation between the charging port and patient accessible parts fails | Use of a damaged or counterfeit USB power source while the pump is worn | Electrical energy | HS_CHG_01 | Mains voltage reaches the enclosure while the pump is connected to the patient | HARM_SHOCK | Electric shock | No detection | 4 | 2 | 5 | 40 | REVIEW | HIGH | engineering estimate | 4 | RCM_CHG_01 Reinforced isolation between the charging port and patient accessible parts [inherent safety by design, verification PASS, credited P1] | REQ_CHG_001 | VER_CHG_001 PASS TR_0119 | 4 | 1 | 5 | 20 | ACCEPTABLE | MEDIUM | 50 | Residual risk acceptable | HDS_002 section 6.3 (DFC_CHG_01); REQ_CHG_001 at HDS_002 section 6.3 | IEC 60601 1 cl 8.5; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2022_002532; MDR_2024_010018; MDR_2025_001459; MDR_2025_006757
RISK_COM_01 | Unauthorized remote bolus command | CMP_COM_01 | Bluetooth Communication Module | Connectivity | FM_COM_01 | Unauthorized remote bolus command is accepted | Weak pairing authentication is exploited by an attacker within radio range | Cybersecurity compromise | HS_COM_01 | A third party triggers insulin delivery without the knowledge of the user | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 2 | 5 | 50 | REVIEW | HIGH | engineering estimate | 4 | RCM_COM_01 Authenticated encrypted pairing with keys generated per session [inherent safety by design, verification PASS, credited P1]; RCM_COM_02 Remote bolus requires confirmation on the pump [protective measure, verification PENDING, credited P0 D0] | REQ_COM_001; REQ_COM_002 | VER_COM_001 PASS TR_0145; VER_COM_002 PENDING TR_0146 | 5 | 1 | 5 | 25 | REVIEW | MEDIUM | 50 | Benefit risk analysis required | UCS_004 section 4.1 (DFC_COM_01); REQ_COM_001 at UCS_004 section 4.1; REQ_COM_002 at UCS_004 section 4.1 | IEC 62304 cl 7.2; IEC 62366 1 cl 5.4; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2021_009144; MDR_2022_001211; MDR_2022_009518; MDR_2025_001547
RISK_COM_02 | Stale glucose data after connection loss | CMP_COM_01 | Bluetooth Communication Module | Connectivity | FM_COM_02 | Connection to the glucose sensor is lost and glucose values stop updating | Radio interference or the transmitter is out of range | Loss of alarm information | HS_COM_02 | The user doses on a stale glucose value that no longer reflects the current level | HARM_HYPO_MODERATE | Moderate hypoglycemia requiring assistance from another person | User observation | 4 | 4 | 4 | 64 | UNACCEPTABLE | HIGH | engineering estimate | 558 | RCM_COM_03 Stale data indicator with rejection of glucose values older than 15 minutes [protective measure, verification PASS, credited P2 D2] | REQ_COM_003 | VER_COM_003 PASS TR_0147 | 4 | 2 | 2 | 16 | REVIEW | MEDIUM | 75 | Benefit risk analysis required | UCS_004 section 4.1 (DFC_COM_02); REQ_COM_003 at UCS_004 section 4.1 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2021_003697; MDR_2022_000796; MDR_2023_006575; MDR_2023_007870; MDR_2024_004758
RISK_COM_03 | Corrupted setting change accepted | CMP_COM_01 | Bluetooth Communication Module | Connectivity | FM_COM_03 | Corrupted data packet is accepted as a valid setting change | Insufficient message integrity checking | Incorrect therapy settings | HS_COM_03 | The basal profile is updated with corrupted values | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 2 | 5 | 50 | REVIEW | HIGH | engineering estimate | 12 | RCM_COM_04 Message authentication code with range validation on all remote setting changes [inherent safety by design, verification PASS, credited P1] | REQ_COM_004 | VER_COM_004 PASS TR_0148 | 5 | 1 | 5 | 25 | REVIEW | MEDIUM | 50 | Benefit risk analysis required | UCS_004 section 4.1 (DFC_COM_03); REQ_COM_004 at UCS_004 section 4.1 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2022_002241; MDR_2023_006506; MDR_2023_008035; MDR_2024_006218; MDR_2024_007075
RISK_DRV_01 | Continuous motor run from a shorted driver stage | CMP_DRV_01 | Stepper Motor Drive Assembly | Drive | FM_DRV_01 | Motor driver stage fails short and the motor runs continuously | Short circuit failure of a transistor in the H bridge driver stage | Overdelivery of insulin | HS_DRV_01 | The pump delivers insulin continuously beyond the programmed dose while the patient is unaware | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 2 | 5 | 50 | REVIEW | HIGH | engineering estimate | 10 | RCM_DRV_01 Independent motor power cutoff by the safety supervisor when encoder counts exceed commanded steps [protective measure, verification PASS, credited P1 D3]; RCM_MCU_03 Independent hourly delivery limit enforced by the safety supervisor [protective measure, verification PASS, credited P1] | REQ_DRV_001; REQ_MCU_003 | VER_DRV_001 PASS TR_0101; VER_MCU_003 PASS TR_0127 | 5 | 1 | 2 | 10 | REVIEW | MEDIUM | 80 | Benefit risk analysis required | HDS_002 section 3.1 (DFC_DRV_01); REQ_DRV_001 at HDS_002 section 3.1; REQ_MCU_003 at SRS_003 section 3.2 | IEC 60601 1 cl 4.2; IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2023_001633; MDR_2023_004907; MDR_2024_000621; MDR_2024_006055; MDR_2024_006239
RISK_DRV_02 | Motor stall without plunger advance | CMP_DRV_01 | Stepper Motor Drive Assembly | Drive | FM_DRV_02 | Motor stalls and the plunger does not advance | Loss of step torque from winding degradation or gear train wear | Underdelivery of insulin | HS_DRV_02 | Programmed basal insulin is not delivered for several hours and no alarm is raised | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | User observation | 4 | 3 | 4 | 48 | REVIEW | HIGH | engineering estimate | 83 | RCM_DRV_02 Encoder based stall detection with a no delivery alarm [protective measure, verification PASS, credited P2 D3] | REQ_DRV_002 | VER_DRV_002 PASS TR_0102 | 4 | 1 | 1 | 4 | ACCEPTABLE | LOW | 91.7 | Residual risk acceptable | HDS_002 section 3.1 (DFC_DRV_02); REQ_DRV_002 at HDS_002 section 3.1 | IEC 60601 1 8 cl 6.1; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2022_007213; MDR_2022_007475; MDR_2023_001462; MDR_2025_006419; MDR_2025_006449
RISK_DRV_03 | Plunger coupling disengagement and free flow | CMP_DRV_02 | Lead Screw and Plunger Coupling | Drive | FM_DRV_03 | Plunger coupling disengages from the lead screw drive nut | Wear or fracture of the coupling latch | Overdelivery of insulin | HS_DRV_03 | Reservoir contents siphon into the patient when the pump is held above the infusion site | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 2 | 5 | 50 | REVIEW | HIGH | engineering estimate | 14 | RCM_DRV_03 Positive locking plunger coupling with retention force margin [inherent safety by design, verification PASS, credited P1] | REQ_DRV_003 | VER_DRV_003 PASS TR_0103 | 5 | 1 | 5 | 25 | REVIEW | MEDIUM | 50 | Benefit risk analysis required | HDS_002 section 3.2 (DFC_DRV_03); REQ_DRV_003 at HDS_002 section 3.2 | ISO 13485 cl 7.3.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2021_001554; MDR_2022_004472; MDR_2023_004158; MDR_2024_004179; MDR_2024_009037
RISK_DRV_04 | Lead screw backlash and dose inaccuracy | CMP_DRV_02 | Lead Screw and Plunger Coupling | Drive | FM_DRV_04 | Lead screw backlash causes inaccurate delivery of small basal increments | Thread wear beyond tolerance over the service life | Underdelivery of insulin | HS_DRV_04 | Delivered basal volume deviates from the programmed volume over an extended period | HARM_HYPER_MILD | Transient hyperglycemia correctable by the user | User observation | 2 | 3 | 4 | 24 | ACCEPTABLE | MEDIUM | engineering estimate | 139 | RCM_DRV_04 Closed loop plunger position correction using encoder feedback [protective measure, verification PASS, credited P2] | REQ_DRV_004 | VER_DRV_004 PASS TR_0104 | 2 | 1 | 4 | 8 | ACCEPTABLE | LOW | 66.7 | Residual risk acceptable | HDS_002 section 3.2 (DFC_DRV_04); REQ_DRV_004 at HDS_002 section 3.2 | ISO 13485 cl 7.3.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_009646; MDR_2022_003941; MDR_2023_006095; MDR_2024_007641; MDR_2025_005550
RISK_ENC_01 | Water ingress | CMP_ENC_01 | Enclosure and Ingress Protection | Enclosure | FM_ENC_01 | Water enters the electronics compartment | Gasket damage after a drop or gasket ageing | Loss of therapy | HS_ENC_01 | Moisture short circuits the electronics and the pump stops delivery without a valid alarm | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 196 | RCM_ENC_01 IPX8 sealed enclosure with drop tested gasket retention [inherent safety by design, verification PASS, credited P2]; RCM_ENC_02 Moisture sensor that triggers an alarm and a transition to the safe state [protective measure, verification PASS, credited D2] | REQ_ENC_001; REQ_ENC_002 | VER_ENC_001 PASS TR_0122; VER_ENC_002 PASS TR_0123 | 4 | 1 | 3 | 12 | ACCEPTABLE | MEDIUM | 80 | Residual risk acceptable | HDS_002 section 7.2 (DFC_ENC_01); REQ_ENC_001 at HDS_002 section 7.2; REQ_ENC_002 at HDS_002 section 7.2 | IEC 60601 1 cl 11.6; IEC 60601 1 cl 4.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_008877; MDR_2022_008750; MDR_2023_004430; MDR_2024_007962; MDR_2025_004161
RISK_ENC_02 | Cracked enclosure | CMP_ENC_01 | Enclosure and Ingress Protection | Enclosure | FM_ENC_02 | Enclosure cracks and exposes a sharp edge | Impact damage from a drop onto a hard surface | Mechanical contact | HS_ENC_02 | A sharp edge of the cracked case rubs against the skin during wear | HARM_LACERATION | Minor skin laceration or irritation | User observation | 1 | 2 | 4 | 8 | ACCEPTABLE | LOW | engineering estimate | 35 | RCM_ENC_03 Impact resistant polymer housing with rounded edge geometry [inherent safety by design, verification PASS, credited P1] | REQ_ENC_003 | VER_ENC_003 PASS TR_0124 | 1 | 1 | 4 | 4 | ACCEPTABLE | LOW | 50 | Residual risk acceptable | HDS_002 section 7.2 (DFC_ENC_02); REQ_ENC_003 at HDS_002 section 7.2 | IEC 60601 1 cl 15.3; IEC 60601 1 cl 9.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_009235; MDR_2023_003456; MDR_2024_005548; MDR_2024_005723; MDR_2025_001567
RISK_INF_01 | Infusion set connector detachment | CMP_INF_01 | Infusion Set Connector | Fluid path | FM_INF_01 | Infusion set connector detaches from the reservoir | Incomplete latch engagement or connector thread wear | Underdelivery of insulin | HS_INF_01 | Insulin is pumped onto skin or clothing instead of into the tissue and no alarm is raised | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | User observation | 4 | 4 | 4 | 64 | UNACCEPTABLE | HIGH | field evidence | 695 | RCM_IFU_04 Training on infusion site checks and routine glucose monitoring [information for safety, verification PASS, credited P1]; RCM_INF_01 Keyed connector with audible and tactile lock confirmation [inherent safety by design, verification PASS, credited P1] | REQ_LBL_004; REQ_INF_001 | VER_LBL_004 PASS TR_0153; VER_INF_001 PASS TR_0112 | 4 | 2 | 4 | 32 | REVIEW | HIGH | 50 | Benefit risk analysis required | HDS_002 section 5.2 (DFC_INF_01); REQ_LBL_004 at UCS_004 section 5.1; REQ_INF_001 at HDS_002 section 5.2 | IEC 62366 1 cl 5.2; IEC 62366 1 cl 5.9; ISO 14971 cl 10.3; ISO 14971 cl 10.4; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2022_007985; MDR_2023_009086; MDR_2024_001438; MDR_2024_003425; MDR_2025_005509
RISK_LBL_01 | Infusion set worn beyond the change interval | CMP_LBL_01 | Labeling and Instructions for Use | Labeling | FM_LBL_01 | Instructions on the infusion set change interval are misunderstood | Ambiguous wording and low readability of the instructions | Use error | HS_LBL_01 | The infusion set is worn beyond the recommended interval | HARM_INFECTION | Infusion site infection | User observation | 3 | 4 | 4 | 48 | REVIEW | HIGH | engineering estimate | 495 | RCM_LBL_01 Infusion site change reminder with configurable interval [protective measure, verification PASS, credited P2]; RCM_LBL_02 Instructions on set change interval validated for comprehension [information for safety, verification PASS, credited P1] | REQ_LBL_007; REQ_LBL_008 | VER_LBL_007 PASS TR_0156; VER_LBL_008 PASS TR_0157 | 3 | 1 | 4 | 12 | ACCEPTABLE | LOW | 75 | Residual risk acceptable | UCS_004 section 5.1 (DFC_LBL_01); REQ_LBL_007 at UCS_004 section 5.1; REQ_LBL_008 at UCS_004 section 5.1 | IEC 62366 1 cl 5.4; IEC 62366 1 cl 5.9; ISO 13485 cl 7.3.7; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_009798; MDR_2022_005446; MDR_2024_002026; MDR_2024_002355; MDR_2025_009997
RISK_MCU_01 | Firmware hang during delivery | CMP_MCU_01 | Main Controller Firmware Platform | Control | FM_MCU_01 | Main controller firmware hangs during delivery | Software deadlock or stack overflow | Loss of therapy | HS_MCU_01 | Delivery halts and the display freezes while no alarm is raised | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 75 | RCM_MCU_01 Hardware watchdog reset with system error alarm from the safety supervisor [protective measure, verification PASS, credited P2 D3] | REQ_MCU_001 | VER_MCU_001 PASS TR_0125 | 4 | 1 | 2 | 8 | ACCEPTABLE | LOW | 86.7 | Residual risk acceptable | SRS_003 section 3.1 (DFC_MCU_01); REQ_MCU_001 at SRS_003 section 3.1 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_007011; MDR_2022_001239; MDR_2023_004511; MDR_2025_001416; MDR_2025_001552
RISK_MCU_02 | Corruption of the active delivery rate | CMP_MCU_01 | Main Controller Firmware Platform | Control | FM_MCU_02 | Memory corruption alters the active delivery rate variable | Stray pointer write or radiation induced bit flip | Overdelivery of insulin | HS_MCU_02 | The pump delivers at a rate many times higher than programmed | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 2 | 5 | 50 | REVIEW | HIGH | engineering estimate | 9 | RCM_MCU_02 Critical variables stored in triplicate with voting and range check before each delivery [inherent safety by design, verification PASS, credited P1]; RCM_MCU_03 Independent hourly delivery limit enforced by the safety supervisor [protective measure, verification PASS, credited P1] | REQ_MCU_002; REQ_MCU_003 | VER_MCU_002 PASS TR_0126; VER_MCU_003 PASS TR_0127 | 5 | 1 | 5 | 25 | REVIEW | MEDIUM | 50 | Benefit risk analysis required | SRS_003 section 3.1 (DFC_MCU_02); REQ_MCU_002 at SRS_003 section 3.1; REQ_MCU_003 at SRS_003 section 3.2 | IEC 60601 1 cl 4.2; IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2021_001796; MDR_2022_002884; MDR_2022_003009; MDR_2023_002347; MDR_2023_003579
RISK_MCU_03 | Silent failure of the safety supervisor | CMP_MCU_02 | Safety Supervisor Processor | Control | FM_MCU_03 | Safety supervisor stops monitoring and the failure remains latent | Supervisor oscillator failure or supervisor firmware fault | Overdelivery of insulin | HS_MCU_03 | A later drive fault is not intercepted because the independent monitor is inactive | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 1 | 5 | 25 | REVIEW | MEDIUM | engineering estimate | 3 | RCM_MCU_04 Mutual heartbeat between main controller and safety supervisor with alarm on loss [protective measure, verification PASS, credited D3] | REQ_MCU_004 | VER_MCU_004 PASS TR_0128 | 5 | 1 | 2 | 10 | REVIEW | MEDIUM | 60 | Benefit risk analysis required | SRS_003 section 3.2 (DFC_MCU_03); REQ_MCU_004 at SRS_003 section 3.2 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2023_009155; MDR_2025_000770; MDR_2025_008385
RISK_MEM_01 | Therapy settings revert to defaults | CMP_MEM_01 | Nonvolatile Settings Memory | Platform | FM_MEM_01 | Therapy settings revert to factory defaults | Flash write interrupted by power loss | Incorrect therapy settings | HS_MEM_01 | The pump resumes delivery with a default basal profile that does not match the prescription | HARM_HYPO_MODERATE | Moderate hypoglycemia requiring assistance from another person | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 81 | RCM_MEM_01 Atomic settings writes with checksum and dual copies, with delivery blocked when settings are invalid [protective measure, verification PASS, credited P2 D2] | REQ_MEM_001 | VER_MEM_001 PASS TR_0139 | 4 | 1 | 3 | 12 | ACCEPTABLE | MEDIUM | 80 | Residual risk acceptable | SRS_003 section 5.1 (DFC_MEM_01); REQ_MEM_001 at SRS_003 section 5.1 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_000727; MDR_2021_004506; MDR_2022_006791; MDR_2022_008810; MDR_2025_001536
RISK_PWR_01 | Battery thermal runaway | CMP_PWR_01 | Lithium Battery Pack | Power | FM_PWR_01 | Battery cell enters thermal runaway | Internal short circuit from a manufacturing defect or mechanical damage | Thermal energy | HS_PWR_01 | The pump surface temperature rises rapidly while the pump is worn against the body | HARM_BURN | Skin burn | No detection | 3 | 2 | 5 | 30 | ACCEPTABLE | MEDIUM | engineering estimate | 13 | RCM_PWR_01 Cell protection circuit with current, voltage and temperature cutoff [protective measure, verification PASS, credited P1]; RCM_PWR_02 Certified cells with incoming inspection and lot traceability [inherent safety by design, verification PASS, credited P1] | REQ_PWR_001; REQ_PWR_002 | VER_PWR_001 PASS TR_0113; VER_PWR_002 PASS TR_0114 | 3 | 1 | 5 | 15 | ACCEPTABLE | LOW | 50 | Residual risk acceptable | HDS_002 section 6.1 (DFC_PWR_01); REQ_PWR_001 at HDS_002 section 6.1; REQ_PWR_002 at HDS_002 section 6.1 | IEC 60601 1 cl 11.1; ISO 13485 cl 7.3.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_003213; MDR_2021_006266; MDR_2023_003533; MDR_2023_007703; MDR_2024_008803
RISK_PWR_02 | Battery depletion without warning | CMP_PWR_01 | Lithium Battery Pack | Power | FM_PWR_02 | Battery depletes suddenly without a low battery warning | Abrupt capacity loss of an aged cell or at low temperature | Loss of therapy | HS_PWR_02 | The pump shuts down during sleep and basal delivery stops for several hours | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 245 | RCM_PWR_03 Staged low battery alarms with reserved alarm capacity [protective measure, verification PASS, credited P1 D3]; RCM_PWR_04 Temperature compensated capacity estimation [inherent safety by design, verification PASS, credited P1] | REQ_PWR_003; REQ_PWR_004 | VER_PWR_003 PASS TR_0115; VER_PWR_004 PASS TR_0116 | 4 | 1 | 2 | 8 | ACCEPTABLE | LOW | 86.7 | Residual risk acceptable | HDS_002 section 6.1 (DFC_PWR_02); REQ_PWR_003 at HDS_002 section 6.2; REQ_PWR_004 at HDS_002 section 6.2 | IEC 60601 1 8 cl 6.1; ISO 13485 cl 7.3.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_004867; MDR_2023_002310; MDR_2023_002736; MDR_2024_003220; MDR_2025_004663
RISK_PWR_03 | Fuel gauge overestimation | CMP_PWR_02 | Battery Fuel Gauge | Power | FM_PWR_03 | Fuel gauge overestimates the remaining battery capacity | Coulomb counter drift without recalibration | Loss of therapy | HS_PWR_03 | The displayed battery level stays normal until the pump shuts down abruptly | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 180 | RCM_PWR_05 Voltage based critical battery threshold independent of the coulomb counter [protective measure, verification PASS, credited P2 D1] | REQ_PWR_005 | VER_PWR_005 PASS TR_0117 | 4 | 1 | 4 | 16 | ACCEPTABLE | MEDIUM | 73.3 | Residual risk acceptable | HDS_002 section 6.2 (DFC_PWR_03); REQ_PWR_005 at HDS_002 section 6.2 | IEC 60601 1 cl 4.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2022_001425; MDR_2022_004026; MDR_2025_002889; MDR_2025_005325; MDR_2025_006312
RISK_RSV_01 | Reservoir seal leak | CMP_RSV_01 | Insulin Reservoir and Seal | Fluid path | FM_RSV_01 | Reservoir seal leaks insulin into the reservoir compartment | O ring compression set or a cracked reservoir barrel | Underdelivery of insulin | HS_RSV_01 | Insulin leaks into the pump instead of reaching the patient while the dose counter still records delivery | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | User observation | 4 | 3 | 4 | 48 | REVIEW | HIGH | engineering estimate | 330 | RCM_IFU_02 Instructions to inspect the reservoir compartment at every reservoir change [information for safety, verification MISSING, credited D0]; RCM_RSV_01 Validated seal compression design with a single use reservoir limit [inherent safety by design, verification PASS, credited P2] | REQ_LBL_002; REQ_RSV_001 | REQ_LBL_002 has no verification record; VER_RSV_001 PASS TR_0109 | 4 | 1 | 4 | 16 | ACCEPTABLE | MEDIUM | 66.7 | Residual risk acceptable | HDS_002 section 5.1 (DFC_RSV_01); REQ_LBL_002 at UCS_004 section 5.1; REQ_RSV_001 at HDS_002 section 5.1 | IEC 62366 1 cl 5.9; ISO 13485 cl 7.3.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_001237; MDR_2022_005970; MDR_2024_000521; MDR_2024_007996; MDR_2024_009519
RISK_RSV_02 | Air bubbles in the reservoir | CMP_RSV_01 | Insulin Reservoir and Seal | Fluid path | FM_RSV_02 | Air bubbles in the reservoir are delivered in place of insulin | Incomplete priming or outgassing of cold insulin as it warms | Underdelivery of insulin | HS_RSV_02 | Air displaces insulin in the infusion line during basal delivery | HARM_HYPER_MILD | Transient hyperglycemia correctable by the user | User observation | 2 | 4 | 4 | 32 | REVIEW | MEDIUM | engineering estimate | 760 | RCM_IFU_03 Illustrated instructions on priming and removal of air bubbles [information for safety, verification PASS, credited P1]; RCM_RSV_02 Guided priming sequence with user confirmation of fluid at the cannula tip [protective measure, verification PASS, credited P1] | REQ_LBL_003; REQ_RSV_002 | VER_LBL_003 PASS TR_0152; VER_RSV_002 PASS TR_0110 | 2 | 2 | 4 | 16 | ACCEPTABLE | LOW | 50 | Residual risk acceptable | HDS_002 section 5.1 (DFC_RSV_02); REQ_LBL_003 at UCS_004 section 5.1; REQ_RSV_002 at HDS_002 section 5.1 | IEC 62366 1 cl 5.4; IEC 62366 1 cl 5.9; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_008538; MDR_2023_005841; MDR_2023_009663; MDR_2024_006502; MDR_2025_000228
RISK_RSV_03 | Loss of sterility of the fluid path | CMP_RSV_01 | Insulin Reservoir and Seal | Fluid path | FM_RSV_03 | Sterile barrier of the reservoir package is breached before use | Pouch seal damage during distribution | Biological contamination | HS_RSV_03 | A contaminated fluid path is connected to the subcutaneous tissue | HARM_INFECTION | Infusion site infection | No detection | 3 | 2 | 5 | 30 | ACCEPTABLE | MEDIUM | engineering estimate | 14 | RCM_RSV_03 Validated sterile barrier packaging with integrity indicator [inherent safety by design, verification PASS, credited P1 D2] | REQ_RSV_003 | VER_RSV_003 PASS TR_0111 | 3 | 1 | 3 | 9 | ACCEPTABLE | LOW | 70 | Residual risk acceptable | HDS_002 section 5.1 (DFC_RSV_03); REQ_RSV_003 at HDS_002 section 5.1 | ISO 13485 cl 7.3.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2022_005420; MDR_2023_001186; MDR_2023_006323; MDR_2024_006910; MDR_2025_007050
RISK_RTC_01 | Clock loses time after battery change | CMP_RTC_01 | Real Time Clock | Platform | FM_RTC_01 | Clock loses time after a battery change | Backup capacitor is depleted during a long battery change | Incorrect therapy settings | HS_RTC_01 | The basal schedule is shifted relative to the true time of day | HARM_HYPER_MILD | Transient hyperglycemia correctable by the user | User observation | 2 | 4 | 4 | 32 | REVIEW | MEDIUM | engineering estimate | 535 | RCM_RTC_01 Clock validity check at power up that forces user confirmation of the time [protective measure, verification PASS, credited P2] | REQ_RTC_001 | VER_RTC_001 PASS TR_0140 | 2 | 2 | 4 | 16 | ACCEPTABLE | LOW | 50 | Residual risk acceptable | SRS_003 section 5.2 (DFC_RTC_01); REQ_RTC_001 at SRS_003 section 5.2 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_004892; MDR_2023_005249; MDR_2024_002885; MDR_2024_005926; MDR_2025_006982
RISK_SNS_01 | Undetected downstream occlusion | CMP_SNS_01 | Occlusion Force Sensor | Sensing | FM_SNS_01 | Force sensor reads low and fails to detect a downstream occlusion | Strain gauge drift or delamination of the load cell | Underdelivery of insulin | HS_SNS_01 | The infusion line is blocked and insulin is not delivered while no occlusion alarm sounds | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 127 | RCM_SNS_01 Force sensor plausibility self test at power on and every 24 hours [protective measure, verification PASS, credited D2]; RCM_SNS_02 Redundant occlusion inference from the motor current signature [protective measure, verification PASS, credited P2 D1] | REQ_SNS_001; REQ_SNS_002 | VER_SNS_001 PASS TR_0105; VER_SNS_002 PASS TR_0106 | 4 | 1 | 2 | 8 | ACCEPTABLE | LOW | 86.7 | Residual risk acceptable | HDS_002 section 4.1 (DFC_SNS_01); REQ_SNS_001 at HDS_002 section 4.1; REQ_SNS_002 at HDS_002 section 4.1 | IEC 60601 1 8 cl 6.1; IEC 60601 1 cl 4.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2022_000604; MDR_2023_008300; MDR_2024_007099; MDR_2025_000939; MDR_2025_005004
RISK_SNS_02 | Unintended bolus on release of an occlusion | CMP_SNS_01 | Occlusion Force Sensor | Sensing | FM_SNS_02 | Occlusion clears suddenly and releases insulin accumulated in the line | Pressure stored behind a kinked cannula is released when the kink straightens | Overdelivery of insulin | HS_SNS_02 | Stored insulin is delivered as an unintended bolus when the occlusion releases | HARM_HYPO_MODERATE | Moderate hypoglycemia requiring assistance from another person | User observation | 4 | 3 | 4 | 48 | REVIEW | HIGH | engineering estimate | 130 | RCM_IFU_01 Instructions directing the user to disconnect before clearing an occlusion [information for safety, verification PASS, credited P1]; RCM_SNS_03 Occlusion alarm threshold that limits stored volume, with plunger retraction on alarm [protective measure, verification PASS, credited P2] | REQ_LBL_001; REQ_SNS_003 | VER_LBL_001 PASS TR_0151; VER_SNS_003 PASS TR_0107 | 4 | 1 | 4 | 16 | ACCEPTABLE | MEDIUM | 66.7 | Residual risk acceptable | HDS_002 section 4.1 (DFC_SNS_02); REQ_LBL_001 at UCS_004 section 5.1; REQ_SNS_003 at HDS_002 section 4.1 | IEC 60601 1 8 cl 6.1; IEC 62366 1 cl 5.9; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_001758; MDR_2021_007195; MDR_2023_000220; MDR_2023_004241; MDR_2023_008009
RISK_SNS_03 | False occlusion alarms | CMP_SNS_01 | Occlusion Force Sensor | Sensing | FM_SNS_03 | Force sensor noise triggers occlusion alarms when no occlusion is present | Mechanical vibration and temperature change shift the force baseline | Loss of therapy | HS_SNS_03 | Delivery stops repeatedly for false occlusion alarms and the user delays restarting delivery | HARM_HYPER_MILD | Transient hyperglycemia correctable by the user | Continuous monitoring with alarm | 2 | 4 | 2 | 16 | REVIEW | MEDIUM | engineering estimate | 599 |  |  |  | 2 | 4 | 2 | 16 | REVIEW | MEDIUM | 0 | Benefit risk analysis required | HDS_002 section 4.1 (DFC_SNS_03) | ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.4 | MDR_2021_000374; MDR_2021_000598; MDR_2021_005816; MDR_2022_009291; MDR_2023_001136
RISK_SNS_04 | Encoder undercount | CMP_SNS_02 | Motor Rotary Encoder | Sensing | FM_SNS_04 | Encoder undercounts motor rotation | Contamination of the optical encoder disc | Overdelivery of insulin | HS_SNS_04 | The controller commands extra motor steps to compensate and delivers more insulin than programmed | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 2 | 5 | 50 | REVIEW | HIGH | engineering estimate | 18 | RCM_SNS_05 Supervisor cross check of encoder counts against commanded steps and motor current [protective measure, verification PASS, credited P1 D3] | REQ_SNS_004 | VER_SNS_004 PASS TR_0108 | 5 | 1 | 2 | 10 | REVIEW | MEDIUM | 80 | Benefit risk analysis required | HDS_002 section 4.2 (DFC_SNS_04); REQ_SNS_004 at HDS_002 section 4.2 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2023_000169; MDR_2023_005284; MDR_2024_000808; MDR_2025_000888; MDR_2025_004733
RISK_SW_01 | Excessive bolus recommendation | CMP_SW_01 | Dose Calculation Software | Therapy software | FM_SW_01 | Bolus calculator recommends an excessive dose | Unit conversion defect between glucose units in the correction calculation | Overdelivery of insulin | HS_SW_01 | The user accepts a recommended bolus several times larger than needed | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | User observation | 5 | 3 | 4 | 60 | UNACCEPTABLE | HIGH | engineering estimate | 97 | RCM_SW_01 Maximum bolus limit and confirmation screen showing the calculation inputs [protective measure, verification PASS, credited P1 D1]; RCM_SW_02 Glucose unit locked at setup with range validation of glucose entries [inherent safety by design, verification PASS, credited P2] | REQ_SW_001; REQ_SW_002; REQ_SW_003 | VER_SW_001 PASS TR_0129; VER_SW_002 PASS TR_0130; VER_SW_003 PASS TR_0131 | 5 | 1 | 3 | 15 | REVIEW | MEDIUM | 75 | Benefit risk analysis required | SRS_003 section 4.1 (DFC_SW_01); REQ_SW_001 at SRS_003 section 4.1; REQ_SW_002 at SRS_003 section 4.1; REQ_SW_003 at SRS_003 section 4.1 | IEC 62304 cl 5.2; IEC 62304 cl 7.2; IEC 62366 1 cl 5.4; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2024_000287; MDR_2024_000771; MDR_2025_001192; MDR_2025_003655; MDR_2025_004906
RISK_SW_02 | Insulin on board not subtracted | CMP_SW_01 | Dose Calculation Software | Therapy software | FM_SW_02 | Insulin on board is not subtracted from a correction bolus | Active insulin time is reset to default after a clock change or software update | Overdelivery of insulin | HS_SW_02 | Repeated correction boluses stack because earlier insulin is ignored | HARM_HYPO_MODERATE | Moderate hypoglycemia requiring assistance from another person | No detection | 4 | 4 | 5 | 80 | UNACCEPTABLE | HIGH | engineering estimate | 398 | RCM_SW_03 Insulin on board persisted with integrity check and shown on the bolus screen [protective measure, verification FAIL, credited P0] | REQ_SW_004; REQ_SW_005 | VER_SW_004 PASS TR_0132; VER_SW_005 FAIL TR_0133 | 4 | 4 | 5 | 80 | UNACCEPTABLE | HIGH | 0 | Design release blocked | SRS_003 section 4.1 (DFC_SW_02); REQ_SW_004 at SRS_003 section 4.1; REQ_SW_005 at SRS_003 section 4.1 | IEC 62304 cl 7.2; IEC 62304 cl 7.4; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2022_004227; MDR_2022_004553; MDR_2023_004546; MDR_2023_009313; MDR_2023_009814
RISK_SW_03 | Basal schedule applied to the wrong time segment | CMP_SW_02 | Delivery Control Software | Therapy software | FM_SW_03 | Basal schedule is applied to the wrong time segment | Scheduler handles a time zone or clock change incorrectly | Incorrect therapy settings | HS_SW_03 | The overnight basal rate is delivered during the day or the daytime rate overnight | HARM_HYPO_MODERATE | Moderate hypoglycemia requiring assistance from another person | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | engineering estimate | 153 | RCM_SW_04 Clock change confirmation with basal schedule review prompt [protective measure, verification PASS, credited P2 D1] | REQ_SW_006 | VER_SW_006 PASS TR_0134 | 4 | 1 | 4 | 16 | ACCEPTABLE | MEDIUM | 73.3 | Residual risk acceptable | SRS_003 section 4.2 (DFC_SW_03); REQ_SW_006 at SRS_003 section 4.2 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_009277; MDR_2022_000020; MDR_2023_006480; MDR_2024_006360; MDR_2025_003573
RISK_SW_04 | Suspended delivery does not resume | CMP_SW_02 | Delivery Control Software | Therapy software | FM_SW_04 | Suspended delivery does not resume when expected | State machine transition defect after alarm acknowledgment | Underdelivery of insulin | HS_SW_04 | The pump remains suspended after the user believes delivery has resumed | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | User observation | 4 | 3 | 4 | 48 | REVIEW | HIGH | engineering estimate | 200 | RCM_SW_05 Persistent suspended indicator with periodic reminder alarm [protective measure, verification PASS, credited P1 D2]; RCM_SW_06 Automatic resume timer that bounds the maximum suspension period [inherent safety by design, verification PASS, credited P1] | REQ_SW_007; REQ_SW_008 | VER_SW_007 PASS TR_0135; VER_SW_008 PASS TR_0136 | 4 | 1 | 2 | 8 | ACCEPTABLE | LOW | 83.3 | Residual risk acceptable | SRS_003 section 4.2 (DFC_SW_04); REQ_SW_007 at SRS_003 section 4.2; REQ_SW_008 at SRS_003 section 4.2 | IEC 60601 1 8 cl 6.1; IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_008498; MDR_2022_003914; MDR_2023_000959; MDR_2023_004861; MDR_2023_005675
RISK_SW_05 | Bolus command executed twice | CMP_SW_02 | Delivery Control Software | Therapy software | FM_SW_05 | Bolus command is executed twice | Duplicate command processing after a communication retry | Overdelivery of insulin | HS_SW_05 | A single bolus request delivers double the intended dose | HARM_HYPO_MODERATE | Moderate hypoglycemia requiring assistance from another person | No detection | 4 | 3 | 5 | 60 | REVIEW | HIGH | field evidence | 52 | RCM_SW_07 Unique command identifiers with idempotent execution and a bolus lockout interval [inherent safety by design, verification PASS, credited P1] | REQ_SW_009 | VER_SW_009 PASS TR_0137 | 4 | 2 | 5 | 40 | REVIEW | HIGH | 33.3 | Benefit risk analysis required | SRS_003 section 4.2 (DFC_SW_05); REQ_SW_009 at SRS_003 section 4.2 | IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 10.4; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2022_007269; MDR_2023_000969; MDR_2023_004082; MDR_2024_007104; MDR_2025_005745
RISK_SW_06 | High priority alarm suppressed | CMP_SW_03 | Alarm Manager Software | Therapy software | FM_SW_06 | High priority alarm is suppressed by lower priority alarms in the queue | Priority inversion in alarm queue handling | Loss of alarm information | HS_SW_06 | An occlusion or no delivery condition is not annunciated | HARM_DKA | Hyperglycemia progressing to diabetic ketoacidosis | No detection | 4 | 2 | 5 | 40 | REVIEW | HIGH | engineering estimate | 15 | RCM_SW_08 Preemptive priority based alarm scheduling [inherent safety by design, verification PASS, credited P1] | REQ_SW_010 | VER_SW_010 PASS TR_0138 | 4 | 1 | 5 | 20 | ACCEPTABLE | MEDIUM | 50 | Residual risk acceptable | SRS_003 section 4.3 (DFC_SW_06); REQ_SW_010 at SRS_003 section 4.3 | IEC 60601 1 8 cl 6.1; IEC 62304 cl 7.2; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_003170; MDR_2021_005530; MDR_2022_004001; MDR_2022_008449; MDR_2023_001776
RISK_UI_01 | Unintended bolus from accidental touch | CMP_UI_01 | Touchscreen and Keypad | User interface | FM_UI_01 | Unintended touch or key activation starts a bolus | Pressure on the touchscreen in a pocket or during sleep | Overdelivery of insulin | HS_UI_01 | A bolus is started without the intent of the user | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | No detection | 5 | 3 | 5 | 75 | UNACCEPTABLE | HIGH | engineering estimate | 110 | RCM_UI_01 Automatic screen lock with unlock gesture and two step bolus confirmation [inherent safety by design, verification PASS, credited P2] | REQ_UI_001 | VER_UI_001 PASS TR_0141 | 5 | 1 | 5 | 25 | REVIEW | MEDIUM | 66.7 | Benefit risk analysis required | UCS_004 section 3.1 (DFC_UI_01); REQ_UI_001 at UCS_004 section 3.1 | IEC 62366 1 cl 5.4; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2021_007820; MDR_2022_001915; MDR_2022_007964; MDR_2024_004903; MDR_2025_001632
RISK_UI_02 | Decimal dose entry error | CMP_UI_01 | Touchscreen and Keypad | User interface | FM_UI_02 | User enters a bolus dose with a misplaced decimal point | Ambiguous numeric entry format leads to a tenfold use error | Use error | HS_UI_02 | A bolus ten times larger than intended is confirmed and delivered | HARM_HYPO_SEVERE | Severe hypoglycemia with seizure, loss of consciousness or death | User observation | 5 | 3 | 4 | 60 | UNACCEPTABLE | HIGH | engineering estimate | 166 | RCM_IFU_05 Training module on bolus entry with knowledge check [information for safety, verification PASS, credited P1]; RCM_UI_02 Fixed decimal dose entry with maximum bolus limit and large dose warning [inherent safety by design, verification PASS, credited P2] | REQ_LBL_005; REQ_SW_001; REQ_UI_002 | VER_LBL_005 PASS TR_0154; VER_SW_001 PASS TR_0129; VER_UI_002 PASS TR_0142 | 5 | 1 | 4 | 20 | REVIEW | MEDIUM | 66.7 | Benefit risk analysis required | UCS_004 section 3.1 (DFC_UI_02); REQ_LBL_005 at UCS_004 section 5.1; REQ_SW_001 at SRS_003 section 4.1; REQ_UI_002 at UCS_004 section 3.1 | IEC 62304 cl 7.2; IEC 62366 1 cl 5.4; IEC 62366 1 cl 5.9; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3; ISO 14971 cl 7.4 | MDR_2022_002778; MDR_2023_000421; MDR_2024_004887; MDR_2024_005801; MDR_2024_008899
RISK_UI_03 | Display unreadable | CMP_UI_01 | Touchscreen and Keypad | User interface | FM_UI_03 | Display fails or becomes unreadable | Backlight failure or a cracked display panel | Loss of alarm information | HS_UI_03 | The user cannot read alarm text or verify bolus entries | HARM_HYPER_SERIOUS | Sustained hyperglycemia requiring medical intervention | User observation | 3 | 3 | 4 | 36 | REVIEW | HIGH | engineering estimate | 133 | RCM_UI_03 Distinct audible alarm patterns per priority that allow response without the display [protective measure, verification PASS, credited P1] | REQ_UI_003 | VER_UI_003 PASS TR_0143 | 3 | 2 | 4 | 24 | ACCEPTABLE | MEDIUM | 33.3 | Residual risk acceptable | UCS_004 section 3.1 (DFC_UI_03); REQ_UI_003 at UCS_004 section 3.1 | IEC 60601 1 8 cl 6.3; ISO 14971 cl 10.3; ISO 14971 cl 5.4; ISO 14971 cl 5.5; ISO 14971 cl 6; ISO 14971 cl 7.1; ISO 14971 cl 7.2; ISO 14971 cl 7.3 | MDR_2021_003069; MDR_2023_007730; MDR_2024_008844; MDR_2025_004256; MDR_2025_007486

## Standards Coverage
Standard | Clause | Topic | Risk items | Requirements | Document sections | Status
IEC 60601 1 | 4.2 | Risk management process for medical electrical equipment | 0 | 5 | HDS_002 section 3.1; HDS_002 section 4.1; HDS_002 section 6.2; HDS_002 section 7.2; SRS_003 section 3.2 | Evidenced
IEC 60601 1 | 8.5 | Separation of parts | 0 | 1 | HDS_002 section 6.3 | Evidenced
IEC 60601 1 | 9.3 | Mechanical hazards associated with surfaces | 0 | 1 | HDS_002 section 7.2 | Evidenced
IEC 60601 1 | 11.1 | Excessive temperatures | 0 | 1 | HDS_002 section 6.1 | Evidenced
IEC 60601 1 | 11.6 | Overflow | 0 | 1 | HDS_002 section 7.2 | Evidenced
IEC 60601 1 | 15.3 | Mechanical strength | 0 | 1 | HDS_002 section 7.2 | Evidenced
IEC 60601 1 8 | 6.1 | Alarm condition | 0 | 6 | HDS_002 section 3.1; HDS_002 section 4.1; HDS_002 section 6.2; SRS_003 section 4.2; SRS_003 section 4.3 | Evidenced
IEC 60601 1 8 | 6.3 | Generation of alarm signals | 0 | 3 | HDS_002 section 7.1; UCS_004 section 3.1 | Evidenced
IEC 62304 | 4.3 | Software safety classification | 0 | 0 | SAS_001 section 5; SRS_003 section 3.1; SRS_003 section 3.2; SRS_003 section 4.1; SRS_003 section 4.2; SRS_003 section 4.3; SRS_003 section 5.1; SRS_003 section 5.2; UCS_004 section 4.1; UCS_004 section 4.2 | Evidenced
IEC 62304 | 5.2 | Software requirements analysis | 0 | 1 | SRS_003 section 4.1 | Evidenced
IEC 62304 | 7.1 | Analysis of software contributing to hazardous situations | 0 | 0 | SRS_003 section 1 | Evidenced
IEC 62304 | 7.2 | Risk control measures | 0 | 21 | HDS_002 section 3.1; HDS_002 section 4.2; SRS_003 section 3.1; SRS_003 section 3.2; SRS_003 section 4.1; SRS_003 section 4.2; SRS_003 section 4.3; SRS_003 section 5.1; SRS_003 section 5.2; UCS_004 section 4.1; UCS_004 section 4.2 | Evidenced
IEC 62304 | 7.3 | Verification of risk control measures | 0 | 0 |  | No evidence in file
IEC 62304 | 7.4 | Risk management of software changes | 0 | 1 | SRS_003 section 4.1 | Evidenced
IEC 62366 1 | 5.1 | Prepare use specification | 0 | 0 |  | No evidence in file
IEC 62366 1 | 5.2 | Identify user interface characteristics related to safety and potential use errors | 0 | 1 | HDS_002 section 5.2 | Evidenced
IEC 62366 1 | 5.3 | Identify known or foreseeable hazards and hazardous situations | 0 | 0 |  | No evidence in file
IEC 62366 1 | 5.4 | Identify and describe hazard related use scenarios | 0 | 6 | HDS_002 section 5.1; SRS_003 section 4.1; UCS_004 section 1; UCS_004 section 3.1; UCS_004 section 4.1; UCS_004 section 5.1 | Evidenced
IEC 62366 1 | 5.9 | Perform summative evaluation of the usability of the user interface | 0 | 8 | UCS_004 section 3.1; UCS_004 section 5.1 | Evidenced
ISO 13485 | 7.3.3 | Design and development inputs | 0 | 6 | HDS_002 section 3.2; HDS_002 section 5.1; HDS_002 section 6.1; HDS_002 section 6.2 | Evidenced
ISO 13485 | 7.3.6 | Design and development verification | 0 | 0 | VVR_005 section 1 | Evidenced
ISO 13485 | 7.3.7 | Design and development validation | 0 | 1 | UCS_004 section 5.1 | Evidenced
ISO 13485 | 8.2.1 | Feedback | 0 | 0 |  | No evidence in file
ISO 14971 | 4.1 | Risk management process | 0 | 0 |  | No evidence in file
ISO 14971 | 4.2 | Management responsibilities | 0 | 0 |  | No evidence in file
ISO 14971 | 4.3 | Competence of personnel | 0 | 0 |  | No evidence in file
ISO 14971 | 4.4 | Risk management plan | 0 | 0 | RMP_006 section 1 | Evidenced
ISO 14971 | 4.5 | Risk management file | 0 | 0 |  | No evidence in file
ISO 14971 | 5.1 | Risk analysis process | 0 | 0 | SAS_001 section 1 | Evidenced
ISO 14971 | 5.2 | Intended use and reasonably foreseeable misuse | 0 | 0 | SAS_001 section 2 | Evidenced
ISO 14971 | 5.3 | Identification of characteristics related to safety | 0 | 0 | SAS_001 section 3 | Evidenced
ISO 14971 | 5.4 | Identification of hazards and hazardous situations | 39 | 0 |  | Evidenced
ISO 14971 | 5.5 | Risk estimation | 39 | 0 |  | Evidenced
ISO 14971 | 6 | Risk evaluation | 39 | 0 |  | Evidenced
ISO 14971 | 7.1 | Risk control option analysis | 38 | 0 | HDS_002 section 2; RMP_006 section 5; SAS_001 section 5; SRS_003 section 2; UCS_004 section 2 | Evidenced
ISO 14971 | 7.2 | Implementation of risk control measures | 38 | 0 | VVR_005 section 1 | Evidenced
ISO 14971 | 7.3 | Residual risk evaluation | 38 | 0 |  | Evidenced
ISO 14971 | 7.4 | Benefit risk analysis | 15 | 0 | RMP_006 section 3 | Evidenced
ISO 14971 | 7.5 | Risks arising from risk control measures | 0 | 0 | RMP_006 section 5 | Evidenced
ISO 14971 | 7.6 | Completeness of risk control | 0 | 0 |  | No evidence in file
ISO 14971 | 8 | Evaluation of overall residual risk | 0 | 0 | RMP_006 section 6 | Evidenced
ISO 14971 | 9 | Risk management review | 0 | 0 | RMP_006 section 6 | Evidenced
ISO 14971 | 10.1 | Production and post production activities | 0 | 0 |  | No evidence in file
ISO 14971 | 10.2 | Information collection | 0 | 0 | RMP_006 section 7 | Evidenced
ISO 14971 | 10.3 | Information review | 39 | 0 | RMP_006 section 7 | Evidenced
ISO 14971 | 10.4 | Actions | 2 | 0 | RMP_006 section 7 | Evidenced

## Gap Report
Level | Rule | Subject | Clause | Finding
BLOCKER | G4 control verification not passed | RCM_SW_03 | ISO 14971 cl 7.2 | RCM_SW_03 for RISK_SW_02 has verification status FAIL, so no credit was given
BLOCKER | G6 unacceptable residual risk | RISK_SW_02 | ISO 14971 cl 7.3 | RISK_SW_02 residual risk is UNACCEPTABLE (severity 4, probability 4)
MAJOR | G1 component without hazard analysis | CMP_ACC_01 | ISO 14971 cl 5.4 | Component CMP_ACC_01 Belt Clip and Holster has no failure mode analysis, yet 260 similar incident reports exist in field data
MAJOR | G2 risk without control | RISK_SNS_03 | ISO 14971 cl 7.1 | RISK_SNS_03 is in the REVIEW region before control and has no risk control measure
MAJOR | G3 control without verification record | RCM_IFU_02 | ISO 14971 cl 7.2 | RCM_IFU_02 for RISK_RSV_01 has an implementing requirement with no verification record, so no credit was given
MAJOR | G4 control verification not passed | RCM_COM_02 | ISO 14971 cl 7.2 | RCM_COM_02 for RISK_COM_01 has verification status PENDING, so no credit was given
MAJOR | G8 field data exceeds engineering estimate | RISK_INF_01 | ISO 14971 cl 10.4 | RISK_INF_01 engineering probability 3 was raised to 4 by 695 similar incident reports
MAJOR | G8 field data exceeds engineering estimate | RISK_SW_05 | ISO 14971 cl 10.4 | RISK_SW_05 engineering probability 2 was raised to 3 by 52 similar incident reports
ACTION | G7 benefit risk analysis required | RISK_CGM_01 | ISO 14971 cl 7.4 | RISK_CGM_01 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_COM_01 | ISO 14971 cl 7.4 | RISK_COM_01 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_COM_02 | ISO 14971 cl 7.4 | RISK_COM_02 residual risk is in the REVIEW region (severity 4, probability 2)
ACTION | G7 benefit risk analysis required | RISK_COM_03 | ISO 14971 cl 7.4 | RISK_COM_03 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_DRV_01 | ISO 14971 cl 7.4 | RISK_DRV_01 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_DRV_03 | ISO 14971 cl 7.4 | RISK_DRV_03 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_INF_01 | ISO 14971 cl 7.4 | RISK_INF_01 residual risk is in the REVIEW region (severity 4, probability 2)
ACTION | G7 benefit risk analysis required | RISK_MCU_02 | ISO 14971 cl 7.4 | RISK_MCU_02 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_MCU_03 | ISO 14971 cl 7.4 | RISK_MCU_03 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_SNS_03 | ISO 14971 cl 7.4 | RISK_SNS_03 residual risk is in the REVIEW region (severity 2, probability 4)
ACTION | G7 benefit risk analysis required | RISK_SNS_04 | ISO 14971 cl 7.4 | RISK_SNS_04 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_SW_01 | ISO 14971 cl 7.4 | RISK_SW_01 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_SW_05 | ISO 14971 cl 7.4 | RISK_SW_05 residual risk is in the REVIEW region (severity 4, probability 2)
ACTION | G7 benefit risk analysis required | RISK_UI_01 | ISO 14971 cl 7.4 | RISK_UI_01 residual risk is in the REVIEW region (severity 5, probability 1)
ACTION | G7 benefit risk analysis required | RISK_UI_02 | ISO 14971 cl 7.4 | RISK_UI_02 residual risk is in the REVIEW region (severity 5, probability 1)